# Contents

Auto ISF — top level

Activity Monitor

AutoISF settings → Libre special settings

AutoISF settings → weight system

AutoISF settings → SMB delivery settings

Notes

Not Yet In This Snapshot (added after build 650)

Related Settings Present In The Export, Outside This Document's Tables

Internal / Non-Setting Keys In This Export

Top | Bottom | TOC

# AutoISF Settings Reference

UKF3426 — every GUI setting under Auto ISF, in menu order, with default values, ranges, and detailed usage notes.

Generated directly from OpenAPSAutoISFPlugin.kt's addPreferenceScreen() (the actual on-device menu structure), the matching key definitions in core/keys/ (defaults and ranges), and cross-referenced against where each setting is actually consumed in the codebase (DetermineBasalAutoISF.kt, related automations) for the "detailed description" column. Scope: this fork's own AutoISF settings tree only, not the whole stock AndroidAPS preference set.

# Auto ISF — top level Back to TOC

| Setting | Default (range) [type] | Detailed description | Current setting (Client) |
| --- | --- | --- | --- |
| MJ Kotlin buttons | on [switch] | Shows a direct "Start MJ" / "Restore MJ" button on the Overview screen instead of requiring the native Automation tab. MJ ("Morning Jog"-style state machine, per this fork's naming) is a coded state that several other automations key off (checkAutomationState("MJ", ...)) to gate profile-switch and boost behaviour. Purely a UI convenience toggle — turning it off does not disable the underlying MJ automations, only the quick-access buttons. | on |
| Steroid Kotlin button | on [switch] | Shows a direct "Steroids ON" button on Overview instead of requiring the native Automation entry. Feeds the steroid-adjustment automations (which progressively raise iobTH/ISF stringency while steroids are active) via the same EventSteroidUserAction path the native automation trigger would use. UI convenience only — the steroid automations themselves are controlled elsewhere. | on |
| Half basal exercise target | 160 (120–200) [number (unit-aware)] | Defines the exercise-sensitivity curve's steepness: at this TT level (with "High temptarget raises sensitivity" on), basal runs at exactly 50%. Other TT levels interpolate along the same curve — e.g. 120 → ~75% basal, 140 → ~60% basal — so lowering this value makes the sensitivity response kick in more aggressively at lower TT levels. | 124 mg/dL / 6.90 mmol/L |

# Activity Monitor Back to TOC

| Setting | Default (range) [type] | Detailed description | Current setting (Client) |
| --- | --- | --- | --- |
| Scale factor for final activity sensitivity | 1.0 (0.0–1.5) [number] | Multiplies activity's default 15–30% sensitivity contribution up or down. Setting to 0.0 disables activity's contribution entirely while leaving inactivity's contribution (below) untouched. | Not in snapshot |
| Scale factor for final inactivity sensitivity | 1.0 (0.0–1.5) [number] | Same idea for inactivity's default 10–20% contribution. 0.0 disables just the inactivity side. | Not in snapshot |

# AutoISF settings → Libre special settings Back to TOC

| Setting | Default (range) [type] | Detailed description | Current setting (Client) |
| --- | --- | --- | --- |
| Offset of Libre calibration versus finger (fslCal_Offset) | 0.0 (-50.0–50.0) [number] | The OFFSET term in this fork's own Libre-vs-reference calibration formula: Libre_calibrated = OFFSET + SLOPE × Libre_raw. Same style of two-point calibration xDrip+ itself does, applied here instead. | 1.35 |
| Slope of Libre calibration versus finger (fslCal_Slope) | 1.0 (0.5–1.5) [number] | The SLOPE term in the same formula. Together with the offset above, this is the LIVE calibration pair actually applied to raw Libre values — it's what "Aging-sensor calibration adjustment" below temporarily overrides during its tiered windows. | 0.74 |
| Libre 1st order exponential smoothing factor | 0.3 (0.1–1.0) [number] | Controls how aggressively raw Libre readings are exponentially smoothed before use, when not relying on xDrip+'s own 1-minute smoothing. 0.0 = maximum smoothing effect, 1.0 = minimum (closest to raw). | Not in snapshot |
| Base Libre cal slope (OldSensorAdj) | 0.72 (0.5–1.5) [number] | The baseline all of Aging-sensor calibration's tiers subtract from: −0.02 (mild, days 2-3 fresh / 11-13 aging), −0.04 (medium, day 1-2 fresh / 13-14 aging), −0.07 (extreme, day 0 fresh / 14-15 aging). Retuning this one value shifts every tier together, rather than needing to update several independent hardcoded numbers. | 0.74 |
| Base Libre cal offset (OldSensorAdj) | 1.4 (-50.0–50.0) [number] | The equivalent baseline for offset: the same tiers add +0.05/+0.10/+0.15 to this value for the mild/medium/extreme windows. | 1.35 |

# AutoISF settings → weight system Back to TOC

| Setting | Default (range) [type] | Detailed description | Current setting (Client) |
| --- | --- | --- | --- |
| Max autoISF ratio (autoISF_max) | 1.0 (1.0–3.0) [number] | Ceiling on how strong the combined adjustment can make ISF. 1.0 disables the strengthening side. Values above 2.5 are explicitly flagged as requiring extreme caution — a large ISF strengthening directly increases how much insulin a given dose delivers. | 2.60 |
| Weight while BG accelerates (bgAccel_ISF_weight) | 0.0 (0.0–1.0) [number] | Strengthens ISF while glucose is accelerating (rate-of-rise increasing), which potentially increases SMB sizing. 0.0 disables this contribution. This is the LIVE value multiple automations write to transiently (during boosts) and restore afterward — see the two settings immediately below. | 0.50 |
| Resting acceleration ISF weight | 0.70 (0.0–1.0) [number] | The value RecentPodOff restores bgAccel_ISF_weight to once its own trigger conditions clear — your normal/baseline acceleration weight, distinct from the live value above which gets overwritten transiently. | 0.70 |
| Boosted acceleration ISF weight | 0.95 (0.0–1.0) [number] | The value OldPod2/RecentPod set bgAccel_ISF_weight to while their own boost conditions are active, and what the AcceUp0.5 automation now also targets. Previously a hardcoded literal in three separate places in the code; now one live, user-tunable setting. | 0.95 |
| Higher ISF-range weight (higher_ISFrange_weight) | 0.0 (0.0–2.0) [number] | Strengthens ISF while glucose is above target, increasing delivery the further above target BG runs. 0.0 disables it. | 0.80 |
| Postprandial ISF weight (pp_ISF_weight) | 0.0 (0.0–0.15) [number] | Strengthens ISF while glucose is actively rising (the "pp" = postprandial signal), increasing SMB sizing during a rise. 0.0 disables it. Like the acceleration weight, this is a LIVE value multiple boost automations overwrite transiently and restore afterward — see the two settings immediately below. | 0.10 |
| Resting postprandial ISF weight | 0.08 (0.0–0.15) [number] | The value boost/recovery automations restore pp_ISF_weight to once a boost window ends. | 0.10 |
| Boosted postprandial ISF weight | 0.15 (0.0–0.15) [number] | The value the fast-rise boost automations (BolusGiven, BolusGivenMild, High6PP, HighOldPod, PodChangeHighPP130, OldPod2, RecentPod) set while actively boosting. Previously hardcoded; now a live setting shared by all of them. | 0.15 |
| dura_ISF weight (dura_ISF_weight) | 0.0 (0.0–3.0) [number] | Strengthens ISF during longer-lasting elevated glucose plateaus ("duration" of a high, not just its level). 0.0 disables it. | 2.50 |
| Resting duration ISF weight | 1.2 (0.0–3.0) [number] | Your normal dura_ISF_weight value. Unlike the acceleration/pp pairs above, nothing currently writes to the live dura_ISF_weight, so this setting isn't yet used as an active restore target — recorded here for when/if that changes. | 2.50 |
| maxIOB threshold percent for SMB (iob_threshold_percent) | 100% (10–100) [number] | SMB delivery is disabled once current IOB exceeds this percentage of your profile's max IOB. 100% (the default) effectively disables this specific throttle — SMB then remains governed only by the hard max-IOB cap itself. | 70% |

# AutoISF settings → SMB delivery settings Back to TOC

| Setting | Default (range) [type] | Detailed description | Current setting (Client) |
| --- | --- | --- | --- |
| FIXED SMB delivery ratio (smb_delivery_ratio) | 0.2 (0.1–1.0) [number] | The share of total calculated insulin requirement that can actually be delivered as a single SMB — a core OpenAPS safety cap, since delivering the full requirement at once would risk stacking. This is also the source value for "UAM Boost insulin req %" internally — that Tier-3 setting has no GUI entry of its own and is instead calculated as this ratio × 100. | 0.15 |
| Simple SMB offset override | off [switch] | When on, replaces the ratio-based SMB derivation entirely with a much simpler rule: SMB is zeroed outright whenever BG is below target + the fixed offset below, full stop. The 8 time-of-day offsets still layer on top either way. Simpler to reason about than tuning the ratio/range pieces above, at the cost of losing their gradation. | on |
| SMB offset (mmol/L above target) | 0.5 (0.0–2.0) [number] | The fixed offset the override above uses. Only has any effect while "Simple SMB offset override" is on. | 0.80 |
| Time-of-day offset, 00–02h | 0.0 (-2.0–2.0) [number] | Added on top of the effective SMB-blocking threshold during 00:00–02:00 only, regardless of which offset mode (ratio-based or simple override) is active. Positive = threshold raised = less SMB allowed in this window; negative = more SMB allowed. | 0.00 |
| Time-of-day offset, 02–04h | 0.0 (-2.0–2.0) [number] | Same mechanism, 02:00–04:00 window. | 0.00 |
| Time-of-day offset, 04–06h | 0.0 (-2.0–2.0) [number] | Same mechanism, 04:00–06:00 window. | 0.00 |
| Time-of-day offset, 06–09h | 0.0 (-2.0–2.0) [number] | Same mechanism, 06:00–09:00 window. | 0.00 |
| Time-of-day offset, 09–12h | 0.0 (-2.0–2.0) [number] | Same mechanism, 09:00–12:00 window. | 0.00 |
| Time-of-day offset, 12–18h | 0.0 (-2.0–2.0) [number] | Same mechanism, 12:00–18:00 window. | 0.00 |
| Time-of-day offset, 18–22h | 0.0 (-2.0–2.0) [number] | Same mechanism, 18:00–22:00 window. | 0.00 |
| Time-of-day offset, 22–00h | 0.0 (-2.0–2.0) [number] | Same mechanism, 22:00–00:00 window (the overnight boundary). | 0.00 |
| Resting SMB delivery ratio | 0.14 (0.1–0.5) [number] | The delivery ratio the boost/recovery automations (DelOff and related) restore to once a boost window ends — a separate, lower baseline from the general SMB delivery ratio above, which those same automations also write into transiently while a boost is active. | 0.15 |
| Enable boost automations (BolusGiven + Mild) | on [switch] | Single master switch for BOTH the BolusGiven family (bg1/bg2/bg3, escalating-severity fast-rise responses) and BolusGivenMild together. Turning it off disables both, not just one. | on |
| Mild boost delivery ratio | 0.20 (0.1–0.5) [number] | Base ratio for BolusGivenMild's own boost — its BGL tiers apply +0.05/+0.02/+0 bumps on top of this value depending on severity. BolusGiven bg3's "strong" tier ratio is this same value +0.03, not configured independently. | 0.23 |
| Enable Tier 3 UAM Boost | off [switch] | Master switch for a separate, newer boost condition in DetermineBasalAutoISF.kt (strong glucose acceleration combined with a positive delta) — distinct from "Enable boost automations" above, which is BolusGiven/BolusGivenMild's own unrelated feature. Also requires VirtualPump as the active pump driver — this Tier 3 logic cannot fire on a real pump build regardless of this toggle, by design, since it's still under real-world evaluation. | Not in snapshot |
| UAM Boost max bolus (boost_max) | 2.5 U (0.1–10.0) [number] | Hard cap on the size of a single UAM Boost SMB. | Not in snapshot |
| UAM Boost max IOB (boostMaxIOB) | 1.0 U (0.1–12.0) [number] | UAM Boost stops firing once current IOB reaches this level, independent of whatever the other trigger conditions say. | Not in snapshot |
| UAM Boost scale (boost_scale) | 1.0 (0.1–3.0) [number] | Multiplier applied to the pre-boost insulin requirement before it's capped at the max bolus above. This is the raw, unscaled preference value — the live SMB size is additionally scaled by the current profile percentage on top of this at dosing time. | Not in snapshot |
| Enable alternative activation of SMB depending on active target being even/odd | off [switch] | An alternate SMB-gating rule: when enabled, SMB is only active while the current target (mmol/L with an even decimal digit, or mg/dL as an even whole number) meets that parity condition — otherwise SMB is inactive. Off = normal behaviour, ungated by target parity. | Not in snapshot |
| Split Bolus | off [switch] | When no profile percentage override is active and a calculated SMB would hit the max-bolus cap, this delivers the remainder as a second part after the split interval below, rather than waiting a full SMB interval for it. | on |
| Show Carb Model Curve | on [switch] | Overlays a dashed, calculated two-compartment carb-absorption curve (90-minute peak, truncated at 240 minutes) on the Carbs Absorption graph, alongside the normal empirical (measured) absorption line — a visual comparison tool, no dosing effect. | on |
| Split Bolus Interval (min) | 7 (1–10) [number] | Minutes between the first and second parts of a split bolus. Only takes effect when Split Bolus is on, profile% = 100%, and the SMB in question is exactly at maxBolus. | 7 |
| TDD Sensitivity (ISF/Basal) | on [switch] | Uses the blended-TDD sensitivity ratio to adjust ISF and basal specifically while Autosens (top-level setting) is off. Disabling this keeps sensitivity neutral (ratio locked to 1.0) instead. | on |
| TDD Factor (max_iob/insulinReq) | on [switch] | Scales both max_iob and the calculated insulin requirement by the same blended-TDD ratio. When off, the fixed fallback value below is used in its place instead. | on |
| TDD Factor fallback value | 1.0 (0.8–1.2) [number] | The fixed TDDfactor used only while "TDD Factor" above is disabled — has zero effect while that's enabled, since the live TDD ratio overrides it every cycle. Range matches the live calculation's own internal clamp. | 1.00 |

# Notes Back to TOC

• "UAM Boost insulin req %" is used internally by Tier 3 UAM Boost's bypass path but has no GUI entry of its own — it is calculated as (FIXED SMB delivery ratio × 100) rather than being independently tunable, so it always tracks that setting.

• Settings marked as a live/transient pair (e.g. “Weight while BG accelerates” plus its “Resting” and “Boosted” counterparts) share one underlying live preference that several automations write to and restore — the “Resting”/“Boosted” settings are the values those automations use, not independent effects of their own.

# Not Yet In This Snapshot (added after build 650) Back to TOC

• Enable Tier 3 UAM Boost — default off, no live value available from this build's export.

• UAM Boost max bolus (boost_max) — default 2.5 U (0.1–10.0), no live value available from this build's export.

• UAM Boost max IOB (% of max_iob) — default 10% (1–100%), no live value available from this build's export.

• UAM Boost scale (boost_scale) — default 1.0 (0.1–3.0), no live value available from this build's export.

# Related Settings Present In The Export, Outside This Document's Tables Back to TOC

Real values from the device, for settings this document's earlier tables don't individually list — included for completeness rather than silently dropped.

• Aging-sensor calibration adjustment (master enable) (autoisf_old_sensor_adj_enabled): true. Referenced in the Libre section's description text but has no row of its own in this document's table.

• Weight while BG decelerates (bgBrake_ISF_weight): 0.25. Sibling of “Weight while BG accelerates”; not currently listed as its own row in this document.

• Enable alternative activation of SMB depending on target parity (Enable alternative activation of SMB always): true. Export label text doesn't exactly match this document's setting name (“...always” vs “...depending on active target being even/odd”) — verify on-device before treating as the same setting.

• Min autoISF ratio / low-BG max ratio (autoISF_min / autoISF_max_low): 0.30 / 1.00. Not present in this document's tables — likely governed elsewhere in the AutoISF weight system code, or scoped out.

• Enable ISF adaptation by glucose behaviour (openapsama_enable_autoISF): true. Plugin-level master switch, outside this document's per-setting tables.

• SMB delivery ratio — BG range / max / min, range extension (openapsama_smb_delivery_ratio_bg_range / _max / _min / smb_max_range_extension): 0.00 / 0.50 / 0.95 / 4.00. Related sibling settings to “FIXED SMB delivery ratio”, not individually tabled in this document.

• High/low temptarget sensitivity toggles (high_temptarget_raises_sensitivity / low_temptarget_lowers_sensitivity): true / true. Stock AAPS sensitivity settings, not part of the AutoISF-specific screens this document covers.

# Internal / Non-Setting Keys In This Export Back to TOC

Present in the raw export file but not GUI settings — internal state, other plugins, or device metadata.

| Preference key | Current value | Why it's not a setting |
| --- | --- | --- |
| autoisf_clean_graph_requested | false | One-shot internal flag, not a setting. |
| autoisf_custom_automations_enabled | true | Not exposed via addPreferenceScreen; read directly by code. |
| autoisf_fslcal_offset_normal / _slope_normal | 1.40 / 0.72 | Internal backup of fslCal_Offset/Slope taken before OldSensorAdj overrides them. |
| autoisf_low_profile_name / autoisf_standard_profile_name | Current Profile / Current ProfileReal | String-type profile-name settings, outside this document's numeric/switch scope. |
| autoisf_low_raw_24_override_active / _armed | false / false | Internal LowRaw24Cal state. Note: _armed no longer exists in the current codebase (removed the same day this export predates). |
| autoisf_low_storage_notified / autoisf_old_pod_notified | false / false | Internal notify-once latches. |
| autoisf_mild_offset_zero_active / autoisf_old_sensor_adj_active | false / false | Internal live-state flags reflecting whether a tier is currently applying. |
| autoisf_sensor_age_code_enabled | true | Related but separate gate from “Aging-sensor calibration adjustment” above; not confirmed to have its own row. |
| automation_state_MJ | MJ3 | MJ state-machine's current value, not a preference. |
| boluswizard_percentage / insulin_oref_peak | 70 / 45 | Belong to the Bolus Wizard / Insulin plugins, not AutoISF. |
| fsl_use_ukf_smoothing / graph5_bgl_only / show_graph5 | false / false / true | Overview double-tap-gesture toggles, not part of the AutoISF preference screens. |
| configuration_* / source / main_phone_snapshot_time | (device metadata) | Export/device metadata, not settings. |